Unitatea de Management al Proiectului ”Prevenirea și Reducerea Poluării din Spațiul Rural în România” - RAPID

Beneficiary: National Administration "Romanian Waters"

Design Brief

Construction of ANAR Regional Laboratory for water quality analysis – Tulcea

Address: Păcii Street, No. 158C, Tulcea Municipality

May 2026

Content

1. General information on investment objective 	2

2. Investment objective identification data 	3

2.1. Information regarding the legal, economic and technical status of the land and/or existing construction, cadastral documentation 	3

2.2. Particularities of the proposed location for achieving the investment objective 	4

a) Brief description of the proposed site (location, land area, plan dimensions): 	4

b) Relationships with neighboring areas, existing accesses and/or possible access routes: 	5

c) Existing pollution sources in the area: 	6

d) Relief features: 	7

e) Level of technical and municipal equipment of the area and possibilities for providing utilities: 	8

f) The existence of any utility networks on site that would require relocation/protection to the extent that they can be identified: 	8

g) Possible easement obligations: 	8

h) Construction conditions determined by the technical condition and the construction system of existing constructions on the sites, on which intervention works will be carried out, as appropriate: 	8

i) Urban planning regulations applicable to the area according to the approved urban planning documentation – PUG / PUZ and the related local urban planning regulation: 	9

j) Existence of historical monuments, architecture or archaeological sites on the site or in the immediately surrounding area: 	10

2.3. Brief description of the proposed investment objective, from a technical and functional point of view 	10

a) Purpose and functions: 	10

b) Specific expected characteristics, parameters and technical data: 	13

c) Level of equipment, finishing and endowment, technical requirements of the construction in accordance with the functional requirements established by technical, heritage and environmental regulations in force: 	14

d) Estimated number of users: 	15

e) Minimum operating time assessed according to the proposed destination/functions: 	15

f) Specific functional needs/requests: 	15

g) Correlation of technical solutions with urban planning, environmental protection and heritage conditions 	17

h) Establishing clear criteria to address the beneficiary's need 	17

2.4. Applicable legislative framework and the impositions resulting from its application 	18

3. ANNEXES 	19

3.1. Annex no. 1 – Indicative functional scenario 	19

3.2. Annex no. 2 – Social and environmental management framework 	19

- information on the investment objective

  - Name of the investment objective

Construction of ANAR Regional Laboratory for water quality analysis – Tulcea

  - Chief Authorizing Officer

- Ministry of Environment, Waters and Forests - the estimated expenses for the development of the PUZ, studies, documentation and expenses for obtaining approvals, agreements and authorizations, technical documentation for obtaining approvals/agreements/authorizations, as well as design services for the Feasibility Study phase and General Cost Estimate are provided from the budget of the Project "Prevention and Reduction of Pollution in Rural Areas in Romania", Component 1 - Modernization of public institutions responsible for pollution control, Subcomponent 1.2 - Strengthening the national capacity for monitoring, prevention and reduction of pollution from agricultural sources in rural areas;

- Ministry of Environment, Waters and Forests - the estimated costs for the DTAC phases, Technical Design and Execution Details, technical quality verification of the project, consulting services, technical assistance, land development and environmental protection, utilities, basic investment, construction site organization, taxes and fees, as well as expenses for technological samples and tests - are proposed to be covered from the budget of the Project "Prevention and Reduction of Pollution in Rural Area in Romania - Additional Financing", for which the approval procedures have been initiated, as well as from other legally established sources.

  - Sub-authorising officer

Not applicable.

  - The beneficiary of the investment

National Administration "Romanian Waters" (ANAR)

  - Design Brief developer

Project Management Unit "Prevention and Reduction of Pollution in Rural Area in Romania", within the Ministry of Environment, Waters and Forests.

- Investment objective identification data

  - Information regarding the legal, economic and technical status of the land and/or existing construction, cadastral documentation

Legal status

The land proposed for the investment is located in Tulcea municipality, on Păcii Street no. 158C. According to the Urbanism Certificate no. 232/12.03.2026, the land is part of the public domain of the Romanian state, under the administration of the National Administration "Romanian Waters" through the Dobrogea-Litoral Water Basin Administration, according to Normative Act no. 1705/29.11.2006.

According to the cadastral documents and the land registry extract no. 34923, the land has a registered area of 1,680 sq m and a measured area of 1,671 sq m, with the use category "yards and buildings". On the land there is the construction C5 - Transformer with an area of 9 sq m built on the ground, which is not under Administration of ANAR.

Currently, the following existing constructions are identified on the site:

C1 – Canton Headquarters: 181 sq_m building footprint

C2 – Warehouse: 112 sq_m building footprint

C3 – Workshop: 64 sq_m building footprint

C4 – Radio Antenna: 8 sq_m building footprint

C5 – Transformer: 9 sq_m building footprint

The land is not subject to any temporary or permanent construction bans and is not encumbered by public utility easements.

Ownership Status

- Owner: Romanian State

- Administration: Dobrogea-Litoral Water Basin Administration

- Current use: yards and buildings and constructions C1, C2, C3, C4, C5

- Legal status: public domain, tabulated

- The site is located in the LM zone - residential with G, G+1, G+2 housing (up to 10 m), LMu2 Subzone - predominantly residential with housing and complementary functions, and P Zone - parks, sports complex and recreation.

- Complementary functions allowed: public institutions, services of general interest, commercial buildings and other services that do not affect the quality of living.

Planning Parameters

- planning parameters: POT = 45%; CUT = 1.3; Maximum number of storeys = 3; maximum cumulative values: POT = 50%; CUT = 1.5; Level = 3,2.

- Maximum permitted height of buildings: 10 m, measured from the lowest point of the land to the roof cornice.

- Alignments and setbacks from the street and plot boundaries: minimum setbacks of 3 m from the side boundaries, minimum 10 m between buildings for the windows of the main rooms.

- Parking spaces: according to HCL no. 201/28.09.2017, maximum 5 spaces per plot; Garages shall be included in the site coverage (POT) calculation but excluded from the gross floor area for (CUT) calculation, in accordance with the applicable urban planning regulations. The final number of parking spaces shall be determined based on the building function, the estimated number of staff and visitors, the operational requirements of the laboratory, the number and type of sampling and service vehicles, accessibility requirements, and the applicable local urban planning and parking regulations. The minimum value indicated in the urban planning documentation shall not be treated as the sole design criterion. The Designer shall substantiate the proposed parking capacity and layout within the Feasibility Study and the related urban planning documentation.

Easements and utilities

- The land has access to a public road and existing utility networks (electricity, water, sewage).

- There are no easements that would prevent the construction of the laboratory.

Requested approvals and documentation

The Urban Planning Certificate requires the following approvals and documents:

- Updated land registry extract and cadastral plan

- Approvals: water supply, sewage, natural gas, sanitation, electricity, fire safety, public health

- Neighbors' notarial agreement – as appropriate

- Notice of opportunity for PUZ

- Zonal Urban Plan approved by H.C.L.

- Topographical study targeted at O.C.P.I.

- Geotechnical study with AF verification

- nZEB compliance report

- Feasibility study

- APM point of view

- Sunlight study – as appropriate

  - Particularities of the proposed location for achieving the investment objective

  - Brief description of the proposed site (location, land area, plan dimensions):

The project site allocated for the construction of the ANAR regional laboratory is located in Tulcea municipality, Păcii Street no. 158C, with an area of 1,680 sq m (measured 1,671 sq m). The land is located within the built-up area of UAT Tulcea, a residential area with G, G+1, G+2 houses and complementary functions.

Land status: public property of the Romanian State, administered by the National Administration of Romanian Waters through ABA Dobrogea-Litoral.

Cadastral number: 34923, according to the CF.

Situation of existing buildings:

| Building code | Function | Area (sqm) | Proposal |
| --- | --- | --- | --- |
| C1 | Canton Headquarters | 181 | Demolition |
| C2 | Shed | 112 | Demolition |
| C3 | Workshop | 64 | Demolition |
| C4 | Radio antenna | 8 | To be retained |
| C5 | Transformer | 9 | To be retained |

  - Relationships with neighboring areas, existing accesses and/or possible access routes:

Boundaries and neighborhoods – the delimitation of space is defined as follows:

- North: 	homes, ~3 m from property line

- South: 	houses, one with a party wall at the property boundary, and outbuildings ~ 3 m from the property boundary

- East: 	Păcii Street

- West: 	homes at ~5 m + existing municipal constructions in the building's yard

The main access to the site is from Păcii Street no. 158C, an existing public road, which ensures circulation for both personnel and supply or sample transport vehicles. It is proposed to maintain a secondary access from the opposite side of the site (Năframei Street), intended for deliveries and technical interventions, as well as the clear organization of internal circulations, by separating pedestrian and car flows, for employees and visitors.

Halfway along the length of the land, there is a level difference of approximately 3 m, which will require the adaptation of access routes and platforms related to the construction, through grading works.

Connection to road infrastructure and utilities will be ensured in accordance with local regulations, without affecting the functions of the neighboring area. The organization of access and traffic will be analyzed and approved by the competent authorities, and it is possible that a traffic permit may be required, depending on the technical solution adopted and the impact on local traffic.

During the Feasibility Study stage, the organization of access will be based on an analysis of traffic flows, with the differentiation, as far as possible, of access for personnel and visitors from logistical access for samples, materials, waste and technical interventions. The solution will aim to limit the impact on residential neighborhoods, ensure access for intervention means and avoid the intersection of clean and potentially contaminated flows.

  - Existing pollution sources in the area:

In Tulcea municipality, air quality is generally classified as good or satisfactory, with normally limited air pollutants (e.g. PM2.5, PM10), according to recent measurements of the Air Quality Index (AQI).

According to available data on air quality in Tulcea municipality, atmospheric pollutant levels generally fall within ranges specific to the urban environment, with variations determined by traffic and weather conditions.

The area of Păcii Street is predominantly urban, without major industrial activities in the immediate vicinity, but with urban traffic and a network of utilities that can generate local variations.

It is useful to continue, in the later design phases, to make a detailed assessment of the quality of the soil and groundwater at the site.

The main potential sources of influence on the site:

| No. | Pollution source type | Distance from the laboratory (Str. Păcii no. 158C) | Main pollutant / type of pollution | Potential damage mode | Observations / recommended measures |
| --- | --- | --- | --- | --- | --- |
| 1 | Urban road traffic (adjacent streets) | opening at Păcii Street – entrance to the site and at Năframei Street – secondary access | NoX, particles (PM10, PM2.5) from traffic | Dust, particle deposition and potential noise; does not directly affect analyses, but may require air filters in the area | Monitoring air quality and ensuring efficient HVAC systems for laboratory and sample reception spaces |
| 2 | Damages and quality of water from the public network used in the laboratory | in the immediate vicinity (problems on Păcii Street) | Temporary turbidity of drinking water after damage | May affect process/cleaning water quality if internal purification systems are not in place | Installation of internal laboratory water supply systems and procedures for testing water before use |
| 3 | Urban residential areas | at variable distances (individual households) | Household waste, nutrients, micropollutants | Low direct impact, but may contribute to overall levels of air pollutants or uncontrolled runoff | Adequate drainage and stormwater management system, regular monitoring |
| 4 | Air pollution | urban general | Common urban pollutants (NO2, CO, PM10 and PM2.5) | Urban air quality may vary seasonally, but is generally good/moderate. | Implementing air filtration for laboratories and monitoring indoor air quality |
| 5 | Potential industrial sources / urban deposits | industrial areas or warehouses > 2 km | Possible chemical/dust emissions | Unlikely direct impact on the laboratory, but requires design assessment | Impact analysis when detailing the project |

  - Relief features:

Halfway along the length of the land, there is a level difference of approximately 3 m, which will require the adaptation of access routes and platforms intended for construction, through earthworks and the arrangement of traffic slopes.

This situation involves the implementation of local land modelling works, including controlled excavation and filling, consolidation of slopes and the arrangement of stable platforms for the location of the building and parking lots, with the provision of an adequate drainage system to prevent water accumulation and erosion phenomena. The technical solutions will be substantiated by a detailed geotechnical study, which will establish the foundation conditions and the necessary measures for the stability of the land.

In general, the lands located in Tulcea municipality, in the area of Păcii Street, present natural unevenness and variable geotechnical conditions, which require the adaptation of foundation solutions and grading works in accordance with specialized studies. Controlled management of rainwater is also required, considering the slightly uneven topography and existing slopes, through collection, evacuation and, where appropriate, retention/infiltration solutions, in accordance with the applicable technical regulations.

Grading works solutions will ensure accessibility for people with disabilities, safe pedestrian routes, roadway access for supply and intervention, as well as stormwater management so that additional flows or uncontrolled leaks are not transferred to neighboring properties.

  - Level of technical and municipal equipment of the area and possibilities for providing utilities:

The location benefits from access to existing municipal networks in the area, with the effective connection possibilities, available capacities and possible expansion/modernization works to be confirmed by utility approvals. For the operation of the laboratory, at least the following categories of utilities and technical infrastructure will be analyzed:

- power supply, including backup solutions for critical equipment, UPS and/or generator;

- drinking water supply and internal systems for producing purified/ultrapure water required for analyses;

- domestic sewage and, where appropriate, separate system for collecting/pretreating technological wastewater resulting from laboratory activities;

- technical gases, as appropriate, with storage/location spaces in compliance with safety requirements;

- IT and communications infrastructure, including secure network for laboratory data management;

- collection, retention and evacuation of rainwater, correlated with the grading works of the land.

  - The existence of any utility networks on site that would require relocation/protection to the extent that they can be identified:

At this stage, there are no confirmed existing utilities on site that would require relocation or protection. The presence, routes and protection/relocation conditions of any water, sewage, electricity, telecommunications or other installations will be clarified through utility permits, topographic surveys, field investigations and related technical documentation.

In the event of identifying networks that cross the site or serve the maintained buildings/installations, the designer will include in the SF solutions for protection, diversion, relocation or maintenance in operation, as appropriate.

  - Possible easement obligations:

At this time, no easements have been identified that would prevent the investment from being carried out, an aspect that will be confirmed by the urban planning documentation and the requested permits.

  - Constructive conditions determined by the technical condition and the constructive system of existing constructions on the sites, on which intervention works will be carried out, as appropriate:

The existing C1–C3 constructions on the site are lightweight constructions, without an essential functional role for the SGA, so it is proposed to demolish them. From a constructive point of view, one of the constructions is made of corrugated sheet metal, and the other two are made of fiberboard (PFL) covered with mesh and plastered. The technical and structural condition of these constructions does not allow their integration into the future investment. The dismantling works will comply with the specific legislation on waste management, including their identification, handling, transport and disposal by authorized operators, in safe conditions for personnel and the environment.

Utility structures C4 and C5 (antenna and transformer) will be maintained and protected during the works, according to the technical plan, to ensure continuity of services and safety of the installations. The legal and technical status of the existing transformer shall be clarified, including its owner, operator, access rights, supplied consumers and applicable safety restrictions. The design documentation shall establish whether the transformer is to be retained, upgraded, protected or relocated and shall include the measures required to ensure safe and continuous power supply during the works.

The design and execution of the investment will be carried out in compliance with the requirements of the Environmental and Social Management Framework set out in Appendix B, including measures to prevent, reduce, monitor and manage environmental and social impacts, in accordance with the applicable provisions of the Loan Agreement, concluded between Romania and the International Bank for Reconstruction and Development (IBRD).

  - Urban planning regulations applicable to the area according to the approved urban planning documentation – PUG / PUZ and the related local urban planning regulation:

According to the Local Urban Planning Regulation related to the General Urban Plan of Tulcea municipality approved by HCL no. 29/27.05.1996; HCL no. 19/25.02.1999 and HCL no. 129/26.05.2011, the building located in Tulcea municipality, 158C Păcii Street, is included in the LMu2 functional area - predominantly residential area with buildings G, G+1, G+2 (up to 10 m), with complementary functions, with the incidence of the P area - green areas, parks, protective curtains, on a setback of approximately 9 m from the street alignment.

The dominant function of the area is urban residential (small individual and collective housing), with a low height regime (G–G+2, maximum 10 m), within an urban structure with complementary functions.

Complementary functions allowed:

- Functions of type ISps - Sub-area for public institutions and services of general interest - constructions for other service provision; ISc - Sub-area for public institutions and services of general interest - Commercial and other constructions that ensure a better quality of housing while maintaining the dominant character LMul - Residential sub-area with buildings G, G+1, G+2 (up to 10m), with urban-type buildings, exclusively residential (housing and complementary functions), but being a transition to LMu2 - Residential area with buildings G, G+1, G+2 (up to 10 m), with urban-type, residential buildings;

- LI type housing - Residential area with buildings with more than 3 floors (over 10 m) that preserve some areas in combination with the IS function - Area for public institutions and services of general interest.

Permitted uses:

- Housing, as the predominant activity on each individual housing plot; the share of housing in the main buildings is at least 75% of the gross floor area of each individual building;

- Activities compatible with habitation, as secondary activities on each plot; these activities can be carried out in the main buildings, in accordance with the normative acts in force and the regulations of the Tulcea Municipality Local Council. The share of space intended for activities compatible with habitation is a maximum of 25% of the SD of each building. Activities complementary to habitation are the following:

• 	Parking or/ and garage for a maximum of 5 cars per plot;

• 	Storage of non-hazardous materials;

• 	Retail trade;

• 	Exercise of liberal professional activities (law, design, medicine, etc.);

• 	Other services.

No temporary building prohibitions applicable to the site have been identified at this stage, according to the urban planning certificate.

The proposed function — regional laboratory for water quality analysis, related to a public institution/service of general interest — will be analyzed and regulated by the Zonal Urban Plan, according to the Urbanism Certificate. The PUZ will establish the location conditions, setbacks, accesses, urban indicators, the relationship with the neighboring residential area and the necessary measures to limit the impact on the neighborhoods.

  - The existence of historical monuments, architecture or archaeological sites on the site or in the immediately surrounding area:

According to the information available at this stage and the provisions of the Urban Planning Certificate, no historical monuments, archaeological sites or protected built areas have been identified that would directly condition the location. If, through the requested permits or during the execution stage, specific conditions are identified, these will be integrated into the technical documentation and execution measures.

  - Brief description of the proposed investment objective, from a technical and functional point of view

The functional data, surfaces, flows, relationships between spaces and technical solutions presented in this Design Brief are indicative and aim to illustrate the complexity level of investment, as well as to define the minimum functional requirements of the beneficiary. The final architectural, installation, compartmentalization and functional organization solutions will be developed and optimized by the designer within the Feasibility Study, in compliance with applicable legislation, technical standards, urban planning conditions and the operation requirements of the beneficiary.

  - Destination and functions:

The functional organization presented in this Design Brief and detailed in the Indicative Functional Scenario (Appendix A) is indicative and establishes the functional relationships, operational flows and minimum requirements necessary for the design of the laboratory. The final architectural, partitioning and internal organization solutions will be developed and optimized by the designer within the Feasibility Study, in compliance with the functional and technical requirements provided for in this Design Brief.

The laboratory building will be used to carry out water quality monitoring and analysis activities for the Danube River Basin, primarily covering the corresponding river basin area Dobrogea-Litoral section, as well as for technical support in cross-border cooperation with the Republic of Moldova and Ukraine.

The building structure will be designed to ensure efficient operational flows, full traceability of samples, personnel safety and functional separation of work areas (reception, preparation, analysis, reporting).

The laboratory will be equipped, depending on the technical solution adopted, with modern equipment and advanced analysis technologies, including inductively coupled plasma mass spectrometry, chromatography, metal analysis workstations and units for the determination of heavy metals and emerging contaminants, including microplastics and pharmaceutical residues, as well as environmental control systems (HVAC) to preserve the integrity of the samples. In addition, the laboratory will have a secure IT infrastructure, with dedicated servers and information systems for sample management and traceability (compatible with LIMS implementation).

The main functions of the laboratory will be the following:

- Sample Reception Room – dedicated space for the receipt, registration and pre-analysis handling of samples, ensuring full sample traceability;

- General chemistry department – for determining the physicochemical parameters of water;

- Biology compartment – for determining biological water quality indicators;

- Heavy metals laboratory – specialized in the analysis of metals and trace elements in water/sediment/biota samples;

- Analysis (equipment) compartments with HVAC control – areas with controlled temperature and humidity, to ensure the accuracy of sensitive analyses (gas chromatography, liquid chromatography, microplastics, etc.);

- Dedicated compartments for sample processing;

- Separate space for balances – dedicated room for precision weighing, isolated from vibrations and temperature fluctuations;

- Hazardous substances storage space – specialized space, according to safety and environmental regulations, for storing hazardous chemical reagents;

- Technical spaces, archive and offices for staff – for carrying out administrative activities, archiving documents and coordinating laboratory activities;

- Conference/training room – for staff training, technical meetings and working sessions with national and international partners;

- Spaces for IT equipment and servers – for the operation of computer systems, data storage and processing, ensuring information security and redundancy;

- Areas for processing, validating and reporting laboratory data;

- Support spaces for field staff – for conducting extended operations and monitoring missions, functionally integrated into the laboratory;

- Logistical and sample handling areas, including temporary storage for samples;

- Areas for handling samples with high biological or chemical risk (e.g. special hoods, controlled access rooms).

The laboratory will operate based on clearly defined operational flows, ensuring full traceability of samples, personnel safety and compliance with the requirements of the SR EN ISO/IEC 17025:2018 standard, as well as with the European regulations applicable in the field of water quality monitoring.

The functional organization will allow the maintenance of the treacebility chain of samples, from reception, unique identification and conditioning/ treatment, to storage, preparation/processing, analysis, validation, reporting and data archiving. The spaces and circuits will be designed to allow access control, recording of critical operations and reducing the risk of contamination, substitution, loss or damage of samples.

The proposed functional organization also enables the integrated performance of water quality analysis activities, ensuring a high level of quality, safety and traceability, in line with the role and responsibilities of a regional laboratory.

Main operational flows:

- Sample flow

- Sample reception will be carried out in a dedicated area, with full traceability and unique identification of each sample;

- Samples will be distributed to specific compartments (biology, general chemistry, heavy metals, organic micropollutants, microplastics) on controlled internal circuits, designed to prevent cross-contamination;

- The analysis results will be collected from the workstations by the information system (LIMS) for validation, storage and reporting of laboratory data;

- The use of a dedicated lift exclusively for the transport of samples, separate from personnel circulation routes, may be considered;

- The analysis results are validated and integrated into the laboratory's IT systems, ensuring complete traceability.

- Staff flow

- To ensure controlled access, distinct circuits will be identified and established for: technical staff; administrative staff; visitors (limited access to non-critical areas, e.g. conference room);

- Access to laboratories is controlled, depending on the authorization level;

- Changing rooms and sanitary spaces will be organized in a way that reduces the risk of external contamination and ensures compliance with biosecurity regulations.

- Material and waste flow

- Hazardous substances and chemical reagents shall be stored in dedicated rooms, equipped with mechanical ventilation and acces control.

- The waste generated will be collected and disposed of according to environmental and safety regulations, through dedicated circuits, separated from the flow of samples and personnel.

- Information and digital flow

- Analytical data are managed through dedicated IT systems, which ensure the unique identification and sample traceability, result validation, archiving and reporting;

- The traceability of each sample will be ensured, and reports can be automatically generated from the IT system for internal and/or external clients.

- Equipment flow and technical maintenance

- The equipment is located in dedicated spaces, with controlled temperature, humidity and ventilation conditions;

- The functional organization shall include dedicated technical areas for HVAC installations, technical gases, IT equipment and support facilities;

- Preventive and corrective maintenance of the equipment is planned to ensure the continued operation of the laboratory and maintenance of equipment performance.

Benefits of internal organization

- Increasing operational efficiency and reducing sample processing times;

- Ensuring compliance with quality and safety standards (SR EN ISO/IEC 17025:2018);

- Improving the quality and reproducibility of analytical results;

- Ensuring full traceability of samples and data;

- The possibility of adapting and expanding the laboratory's functions according to technological developments and European requirements.

The organization of the spaces will respect the principles of separation of clean and potentially contaminated flows, as well as laboratory security requirements. At the same time, it allows the laboratory to function as an integrated system, in which reception, analysis and reporting activities are efficiently correlated, ensuring a high level of quality, safety and traceability.

Indicative functional diagrams of operational flows:

The functional diagrams below are indicative and aim to illustrate the principles of operational organization, functional relationships and the degree of complexity specific to the laboratory activities. The final architectural, partitioning and circuit organization solutions will be developed and optimized by the designer within the Feasibility Study.

  - Expected specific features, parameters and technical data:

The ANAR Tulcea regional laboratory will be sized to meet the needs of water quality analysis in the Danube Basin - Dobrogea - Litoral, as well as to support cross-border cooperation activities with the Republic of Moldova and Ukraine. The laboratory capacity will be established depending on the estimated volume of samples and the type of analyses, with the possibility of expansion in the medium and long term.

Sizing criteria

- types of analyses and laboratory profile: chemical analyses, heavy metal determinations, analyses for emerging pollutants;

- systems and infrastructure: controlled HVAC systems for critical areas, purified water systems, redundant power supply, safety and environmental protection systems;

- other particularities: administrative functions, training rooms, IT and server spaces, storage of hazardous substances;

- location and terrain: stable terrain in terms of geotechnical risk; any foundation measures will be established based on the geotechnical study.

Expected characteristics, parameters and technical data for the laboratory

- the estimated capacity will be established and substantiated within the SF, depending on the operational needs and types of analyses;

- total land area: 1,680 m², with trapezoidal plan: large base = ~72 m, small base = ~67 m, height = ~21 m; street opening = ~23 m;

- built-up area (SC): approximately 700 m²;

- gross floor area (SD): approximately 2,100 m²;

- site covering (POT): 700/1,680 = 41.67%;

- estimated height regime: ground floor + 2 floors (maximum 10 m), with the possibility of creating a landscaped terrace, technical volumes and/or a secluded level, depending on the architectural solution and the applicable urban planning regulations;

- utilities: drinking and technological water, electricity with back-up system, technical gases, as appropriate, sewage for wastewater, equipped with collection and pre-treatment systems adapted to the specifics of laboratory activities, in accordance with environmental regulations as well as communications and data transfer infrastructure (internet, voice and data network, optical fiber), necessary for the operation of equipment, IT systems and digital laboratory platforms.

  - The level of equipment, finishing and endowment, technical requirements of the construction in accordance with the functional requirements established by technical, heritage and environmental regulations in force :

The laboratory design will aim to meet the conditions necessary to carry out specific water quality monitoring and analysis activities, in accordance with the requirements of the SR EN ISO/IEC 17025:2018 standard, the applicable technical regulations and the norms regarding biological safety, environmental protection and energy efficiency.

The technical and constructive solutions will aim to:

- functional separation of sample, personnel, material and waste flows;

- ensuring controlled temperature, humidity and ventilation conditions in laboratory spaces;

- flexibility and possibility of subsequent adaptation of spaces and facilities;

- operational safety and reduction of contamination risks;

- ensuring universal accessibility and barrier-free access for persons with disabilities and persons with reduced mobility, in accordance with the applicable legislation;

- energy efficiency and compliance with nZEB requirements.

The building may include but not be limited to, as appropriate, the following systems and equipment:

- HVAC installations dedicated to laboratory spaces, including filtration systems and environmental parameter control;

- electrical and low-voltage installations, with backup sources for critical equipment;

- sanitary installations and technological water networks;

- systems for technical gases and controlled exhausts;

- fire detection and extinguishing systems, access control and video surveillance;

- IT and communications infrastructure for evidence management and traceability;

- equipment and furniture specific to analysis laboratories for monitoring water quality.

Ventilation and air conditioning systems will be differentiated by functional areas, taking into account the requirements of temperature, humidity, pressure, local exhaust, chemical hoods, reagent storage and personnel protection. Air recirculation from areas with chemical or biological risk will be avoided, where technical regulations and risk analysis require controlled exhaust.

For laboratory spaces, washable finishes, resistant to chemicals, easy to sanitize and compatible with safety and operation requirements will be considered. Floors, walls, countertops and technical furniture will be selected depending on the type of activities carried out, the chemical/biological risks and the need to maintain controlled working conditions.

The investment will also take into account:

- arrangement of technical and administrative spaces;

- organization of access and functional circulations;

- creating the platforms and external arrangements necessary for the operation of the laboratory, including dedicated outdoor storage areas for technical gas cylinders;

- integrating energy efficiency solutions and using renewable energy sources.

The detailing of technical solutions, installations, finishes and equipment will be carried out during the Feasibility Study stage, based on the functional requirements established by this Design Brief.

  - Estimated number of users :

The building will be designed for an estimated 15–20 permanent users, representing technical, administrative and support staff. The sizing of common areas, toilets, changing rooms, training/conference room, parking and evacuation routes will also take into account the temporary presence of visitors, auditors, maintenance staff, institutional collaborators and participants in training activities. The final number of users and the sizing of the spaces will be analyzed and optimized during the Feasibility Study stage.

  - Minimum operating time estimated according to the proposed destination/functions:

Based on the lifespan of buildings with similar purpose and the use cycle of high-performance laboratory equipment, a minimum operating life of:

- 40-50 years for construction;

- 5–15 years for equipment and facilities, depending on the type and intensity of use.

Maintaining these parameters is conditional on the implementation of a continuous program of maintenance, modernization and phased replacement of equipment, as well as permanent adaptation to the requirements of the SR EN ISO/IEC 17025:2018 standard and to technological developments in the field of analyses for monitoring water quality.

  - Needs /requests:

The operation of the ANAR regional laboratory for water quality analysis involves the provision of specialized facilities, spaces and infrastructures, adapted to monitoring activities, physico-chemical, biological and instrumental analysis of water quality, including the determination of traditional and emerging pollutants, as appropriate. The functional organization and facilities will be correlated with the requirements of the SR EN ISO/IEC 17025:2018 standard, with the applicable legislation in the field of water monitoring, with occupational health and safety norms and with environmental protection requirements.

Depending on the proposed indicators, the facilities will be provided in order to allow the application of internal quality control procedures to ensure the validity of results, traceability of measurements, calibration/verification of equipment, participation in interlaboratory comparison schemes and maintenance of the conditions necessary for accreditation according to SR EN ISO/IEC 17025:2018. The following categories of equipment and facilities are considered, as guidance, within the investment:

Laboratory equipment, indicative

- spectrophotometers, liquid and gas chromatographs with different detectorsICP-MS/ICP-OES equipment for metal determination, microscopes and equipment specific to the biology department, other equipment for chemistry;

- equipment for sample preparation and processing: extraction, evaporation  and concentration system centrifuges, shakers, filtration systems, digestion/mineralization, automatic pipettes and other auxiliary equipment;

- measurement and control equipment: analytical balances, multiparameters, turbidimeters, TOC/TN analyzers, temperature, humidity and other relevant parameter monitoring systems, etc.;

- equipment for storing samples: refrigerators, freezers, ultra-low temperature freezers, cold rooms, cabinets for samples and counter-samples, storage cabinets for chemicals, reagents and solvents, etc..

Special installations and equipment

- Controlled exhaust chemical hoods;

- Dedicated HVAC systems, with differential pressure and contamination control, adapted to chemical and biological work areas;

- Networks for technical gases (He, N₂, H₂ etc.), including nitrogen generators;

- Ultra-pure water systems;

- Dedicated and ventilated spaces for storing hazardous substances (solvents, chemical and biological reagents);

- Specialized areas for the collection and management of chemical and biological waste;

- Environmental condition monitoring systems (temperature, humidity, pressure, gas detection).

- Uninterruptible Power Supply (UPS) systems to ensure the continuous and stable power supply required for the operation of analytical equipment.

Furniture and internal infrastructure

- Technical laboratory furniture:

  - laboratory benches, chemical fume hoods and other laboratory furnitureetc;

  - cabinets for reagents and solvents;

  - spaces dedicated to sensitive equipment;

- Organized spaces for:

  - storage of samples;

  - archiving;

- Technical lift for transporting samples and materials between levels.

IT systems and laboratory software

- Dedicated workstations and servers for operating equipment and processing data;

- Software specific to data collection for analytical equipment (chromatography, spectrometry, etc.);

- Information systems for laboratory data management, sample traceability and reporting;

- Secure connectivity for data;

- Scalable IT infrastructure, compatible with the subsequent implementation of an integrated LIMS system.

Safety and personnel protection equipment

- Detection and alarm systems for gases and hazardous substances;

- Emergency showers and eyewash equipment;

- Personal protective equipment (PPE) adapted to chemical and biological activities;

- Access control systems for functional areas, depending on the risk level;

- Procedures and infrastructure for ensuring biosecurity and chemical security.

Other functional requirements specific to the ANAR regional laboratory

- Spaces for training and knowledge transfer (conference room);

- Modular organization of spaces, allowing adaptation to future technological requirements;

- Circuit separation:

  - samples / personnel / visitors;

  - clean areas / areas with chemical or biological risk.

Preliminary Room Data Sheets shall be prepared for the main laboratory, technical and support spaces. Each sheet shall identify the room function, estimated number of users, main equipment categories, indicative area, general environmental conditions, special ventilation requirements, main utility needs, relevant structural constraints, safety requirements and functional relationships with adjacent spaces. Detailed equipment data, final system capacities, exact utility connections, installation layouts, outlet locations, finishes and construction details shall be developed during further Design stage.

  - Correlation of technical solutions with urban planning, environmental protection and heritage conditions

The proposed technical solutions will be correlated with the urban planning regulations applicable to the site in Tulcea municipality, including the provisions of the Urban Planning Certificate, the approved urban planning documentation and the related local regulations. The design will aim to comply with urban planning indicators (POT, CUT, height regime, setbacks, accesses and circulations), as well as the functional and architectural integration of the objective in the existing area.

At the same time, environmental protection requirements will be taken into account, including controlled management of chemical and biological waste, energy efficiency, emission reduction and the use of solutions compatible with nZEB principles. Design solutions will aim to limit the impact on the environment and ensure biosecurity and safety conditions specific to laboratory activities.

In the event of identification of conditions regarding heritage protection or other site-specific restrictions, these will be analyzed and integrated into the technical solutions developed within the Feasibility Study and subsequent design stages.

  - Establishing clear criteria to address the beneficiary's needs

The Design Brief establishes the functional, technical and operational requirements necessary to create a regional laboratory designed as a strategic infrastructure for the monitoring and protection of water resources within the Dobrogea-Litoral and Lower Danube ABA.

In order to substantiate the technical solutions and ensure a unified approach in the subsequent design stages, the following main criteria are taken into account:

- ensuring separate functional flows for samples, personnel, materials, waste;

- compliance with biosecurity, safety and environmental protection requirements;

- flexibility and the possibility of adapting spaces and infrastructure to the evolution of technologies and operational requirements;

- ensuring the necessary conditions for the operation of high-performance laboratory equipment;

- energy efficiency and integration of sustainable solutions;

- ensuring universal accessibility and inclusive design principles, in accordance with the applicable legislation and the functional requirements of the investment;

- interoperability of IT systems and ensuring full traceability of samples and results;

- ensuring the necessary conditions for accreditation and operation according to / SR EN ISO/IEC 17025:2018 requirements;

- optimizing long-term operating, maintenance and maintenance costs.

These criteria will form the basis for the development and evaluation of the solutions proposed in the Feasibility Study and subsequent design stages.

  - The applicable legislative framework and the obligations resulting from its application

The investment regarding the creation of the ANAR Regional Laboratory in Tulcea will be developed in compliance with applicable national and European legislation in the field of construction, laboratory infrastructure, environmental protection, occupational health and safety, energy efficiency and water quality monitoring activities.

When developing the technical and economic documentation and the project, the following will be taken into account, mainly:

- GD No. 907/2016 on the stages of development and the framework content of technical and economic documentation, with subsequent amendments and supplements;

- Law No. 10/1995 on quality in construction, with subsequent amendments and supplements;

- Law No. 50/1991 on the authorization of construction works, with subsequent amendments and supplements;

- Law No. 448/2006 on the protection and promotion of the rights of persons with disabilities and the related technical regulations governing accessibility of public buildings, as amended;

- building energy performance regulations and nZEB requirements;

- fire safety regulations;

- regulations on environmental protection and waste management;

- SR EN ISO/IEC 17025:2018 standard;

- European and national legislation in the field of water quality monitoring;

- regulations regarding data protection and information systems security, to the extent that the laboratory will use dedicated digital systems/LIMS.

Given the clearly defined nature of the investment, the unique location and the specific technical requirements, the investment can be substantiated directly through the Feasibility Study, in accordance with the provisions of GD no. 907/2016. In subsequent stages, the following will be developed, as appropriate: PUZ approved and approved according to applicable legislation; topographic study covered by OCPI; geotechnical study verified according to applicable requirements; technical and technological studies necessary to define laboratory functions; documentation for obtaining the approvals / agreements / authorizations requested through the Urbanism Certificate; environmental documentation; nZEB compliance report; fire safety scenario, as appropriate; technical documentation related to subsequent design stages.

- APPENDIX

  - Appendix A – Indicative functional scenario

  - Appendix B – Social and environmental management framework